
Micro Electro Mechanical System (MEMS) for the measurement of density and viscosity

A.R.H. Goodwin1, E.P. Donzier1, O. Vancauwenberghe1, and S. Jacobs2

1 Schlumberger-Doll Research, Ridgefield, Connecticut, USA 

2 Schlumberger Oilphase DBR, Edmonton, Alberta, Canada
Measurements of the density and viscosity of gases are required to determine the value of the produced fluid and production strategy.  In this work, a Micro Electrical Mechanical System (MEMS) has been developed to determine both density and viscosity of fluids.  This device, shown in figure 1, is based on a vibrating plate, with dimensions on the order of 1 mm and a mass of about 0.3 mg, clamped along one edge. The plate is set in motion when an ac current is passed through the coil, shown in figure 1, atop the plate in the presence of a magnetic field.  At resonance the plate motion is determined with a strain gauge, shown in figure 1.  The resonance frequency, under vacuum, of the first bending mode is about 12 kHz, at a temperature of 298 K, with a quality factor of about 2900.  Measurements of the resonance frequency and quality factor of the vibrating plate were combined with mechanical properties of the plate to determine, with a hydrodynamic working equation, the density and viscosity of the fluid surrounding it at temperature and pressure.  

[image: image1.jpg]By
SEHLUMBERASR |

H§ PRATENT
PEHDING





Figure 1.  Photograph of the upper surface of the MEMS for measurement of density and viscosity.  The active element, which is a rectangular plate, is below the coil and shown at left. 

Densities, in the range (1 to 1300) kg(m-3, and viscosities, in the range (10 to 50000) Pa s, were determined with the vibrating plate for argon, methane, nitrogen, natural gas, octane, toluene, other reservoir fluids and brine.  These measurements, when compared with either accepted literature values or results obtained with experimental techniques that utilise different principles, and thus have quite different sources of systematic error, were found to lie within (0.5 per cent for density and (5 per cent for viscosity of the literature values with a confidence of 95 per cent. 

Magnetic 


field





Coil





Resistance thermometer





Strain gauge















